Protocol for Rehab tools experiment.	*these protocols assume that the network is functional and that

NeatTools software is loaded and up to date on all relevant

computers.

Set up the physical system

Private Satellite network.

Be sure a receive computer with neattools and grecieve.ntl loaded is set up

at the location.

Connect the computer in Hawaii to the Ethernet bridge on the satellite net

Note the IP address of this computer..

Help the simulated patient to don the knee brace with goiniometer.

Put on the thigh and calf pads snuggly with the included Velcro.

Add the brace, be sure that the flexion point of the brace is parallel

with the knee.

Strap the brace on snuggly.

Connect the 1/8 “ stereo plug on the brace through an extension cord to the

receptical for channel 1 on the TNG-3B box.

Connect a 9 pin D serial cable from the TNG-3B to the COM-1 port on the computer

Comercial Internet LAN connection.

Same setup as above except the Ethernet bridge in Hawaii is connected to the

T1 commercial internet connection via Ethernet bridge and air LAN.

Analog Telephone PPP connection.

Be sure that both send and receive computers have modems and Dialup networking

Adapters, with TCP/IP protocol, and Dial up server installed.

Connect the two computers through Dial up server, not the IP address of the send

Computer; follow the setup listed under PCN for other details.

HAM radio connection.

Connect each computer through a TNC to a HAM radio, allow the HAM engineers

to establish IP connection between machines, otherwise same as above.

Run the Software/Calibrate if necessary

At the receive computer run grecieve.ntl.

At the send computer run gsend.ntl, push the purple ‘on’ button.  Have the simulated client move

his/her leg (and thus the goiniometer through a full range of motion note the changing angles on

the screen.

If this goiniometer has not been calibrated, do so now by selecting the calibrate button

and allowing the goiniometer to go through a full range of motion, then deselect the

calibrate button.  Now edit the full range maximum and minimum values indicated to

the right of the calibrate button.  Set the clients knee to exactly 90 degrees (measured in

real space not the computer, and push the set 3rd point button on the interface so that the

computer angle reads 90.

Simultaneously transmit and record data

On the recieve computer, edit the Socket Client, input the IP address of the send computer

in the place of “localhost” on the module properties.

On the receive computer push the green button to allow communication, the Socket Client

module’s red indicator light should turn yellow, indicating that communications are open

but not running.

On the send computer, push the Green button to start communication, the Socket Client

Module’s indicator light should turn red on both computers indicating that communication is

happening.

When you see data coming through on the receive computer edit the recoring filename to

a unique identifier, this should be the same on both computers, push the blue button to begin

recording, make sure that both computers are recording data.

Have the simulated client move discretely through ten different positions, record them

manually for back up.

Repeat with multiple simulated clients.

Compare recorded data files

After completeing all the trials transmit the datafiles saved on the sending computer

through ftp or windows calls to the receive computer.  Note the amount of time it takes to

transer these files over each system.

Compare the recorded data files, note any differences,